# CHRONOS-Rb

Compact High-precision Rubidium Oscillator Network Operating System

Hardware Design & Assembly Guide

Raspberry Pi Pico 2-W NTP/PTP Server

Synchronized to FE-5680A Rubidium Frequency Standard

Version 1.0

Open Source Hardware Project

# 1. Project Overview

CHRONOS-Rb is a high-precision NTP/PTP time server based on the Raspberry Pi Pico 2-W microcontroller, disciplined by an FE-5680A rubidium atomic frequency standard. This combination provides Stratum-1 quality time distribution over WiFi with sub-microsecond stability.

## 1.1 Key Features

- Stratum-1 NTP server with rubidium reference (10⁻¹¹ stability)

- IEEE 1588 PTP support for precision timing applications

- WiFi connectivity via Pico 2-W's CYW43439 chip

- Web-based status and configuration interface

- PIO-based precision timing for sub-microsecond accuracy

- Automatic holdover during reference signal loss

- Five interval pulse outputs (0.5s, 1s, 6s, 30s, 60s) synchronized to atomic reference

# 2. FE-5680A Rubidium Oscillator

The FE-5680A is a compact rubidium atomic frequency standard originally designed for telecommunications. It provides an extremely stable 10MHz output signal locked to the hyperfine transition of rubidium-87 atoms.

## 2.1 Specifications

| Parameter | Value | Notes |
| --- | --- | --- |
| Output Frequency | 10 MHz | Sine wave |
| Output Level | ~1 Vpp | Into 50Ω |
| Short-term Stability | < 3×10⁻¹¹ | At 1 second |
| Long-term Stability | < 5×10⁻¹² | At 1 day |
| 1PPS Output | TTL compatible | ~100µs pulse |
| Supply (+15V) | 15V DC | Physics heater |
| Supply (+5V) | 5V DC | Electronics |
| Warmup Current | ~2A @ 15V | First 3-5 min |
| Operating Current | ~0.7A @ 15V | After warmup |
| Warmup Time | 3-5 minutes | To lock |
| Lock Indicator | Active LOW | Open collector |

## 2.2 FE-5680A Pinout

| Pin | Signal | Description |
| --- | --- | --- |
| 1 | +15V | Main power input |
| 2 | GND | Ground |
| 3 | +5V | Logic power |
| 5 | 10MHz OUT | Sine wave (1Vpp) |
| 7 | 1PPS OUT | Pulse per second |
| 9 | LOCK | Lock status (active LOW) |

# 3. Signal Conditioning Circuit

The FE-5680A outputs a 10MHz sine wave at approximately 1Vpp, which must be converted to a 3.3V digital signal for the Pico 2-W. The 1PPS signal also needs level shifting from 5V TTL to 3.3V LVCMOS.

## 3.1 10MHz Sine to Square Converter

A high-speed comparator converts the 10MHz sine wave to a clean 3.3V square wave. The LT1016 or MAX999 are recommended for their fast propagation delay (<10ns).

The circuit uses AC coupling (C1) to remove DC offset. R3 sets the threshold at ~1.65V. R1 provides pull-up for the open-drain output.

## 3.2 1PPS Level Shifter

A simple voltage divider converts the 5V TTL 1PPS signal to 3.3V, with a Schottky diode for ESD protection.

# 4. Power Supply Design

The FE-5680A requires +15V (physics heater) and +5V (electronics). The Pico needs 3.3V via its onboard regulator from 5V USB or VSYS.

## 4.1 Power Requirements

| Rail | Voltage | Warmup | Running |
| --- | --- | --- | --- |
| +15V | 15V DC | 2.0A | 0.7A |
| +5V | 5V DC | 0.3A | 0.2A |
| +3.3V | 3.3V DC | 0.3A | 0.15A |

## 4.2 Power Configuration

## 4.3 Safety Notes

- WARNING: FE-5680A physics package runs HOT. Ensure ventilation.

- Add a 3A fuse in series with +15V supply.

- Never connect/disconnect signals while powered.

# 5. Complete Wiring Guide

## 5.1 GPIO Assignments

| GPIO | Function | Connection |
| --- | --- | --- |
| GP2 | 1PPS Input | From level shifter |
| GP3 | 10MHz Input | From comparator |
| GP4 | Lock Status | From FE-5680A pin 9 |
| GP6 | LED Sync | Green LED |
| GP7 | LED Network | Blue LED |
| GP8 | LED Activity | Yellow LED |
| GP9 | LED Error | Red LED |
| GP10 | Debug PPS | Test output |
| GP12/13 | I2C | Optional OLED |
| GP14 | Pulse 0.5s | 500ms interval |
| GP15 | Pulse 1s | 1 second interval |
| GP16 | Pulse 6s | 6 second interval |
| GP17 | Pulse 30s | 30 second interval |
| GP18 | Pulse 60s | 60 second interval |

## 5.2 System Schematic

## 5.3 Status LED Circuit

Four LEDs provide visual feedback. Each is driven through a 330Ω current-limiting resistor.

## 5.4 Interval Pulse Outputs

CHRONOS-Rb provides five interval pulse outputs synchronized to the rubidium 1PPS reference. These outputs generate 10ms active-high pulses at precise intervals, useful for triggering external equipment, calibration, or timing distribution.

| Output | GPIO | Interval | Use Case |
| --- | --- | --- | --- |
| 500ms | GP14 | 0.5 seconds | High-rate timing, servo sync |
| 1s | GP15 | 1 second | PPS distribution |
| 6s | GP16 | 6 seconds | GPS-compatible timing |
| 30s | GP17 | 30 seconds | Calibration triggers |
| 60s | GP18 | 60 seconds | Minute markers |

All outputs are phase-locked to the atomic 1PPS reference and maintain sub-microsecond accuracy. The 0.5s output pulses twice per second (at 0ms and 500ms after each PPS edge). Use external buffers (e.g., 74LVC1G125) if driving cables or loads requiring more than 12mA.

# 6. Bill of Materials

## 6.1 Main Components

| Qty | Description | Part Number | Notes |
| --- | --- | --- | --- |
| 1 | Raspberry Pi Pico 2-W | SC1623 | With WiFi |
| 1 | FE-5680A Rubidium | FE-5680A | Surplus/eBay |
| 1 | 15V 3A PSU | - | Min 45W |
| 1 | LT1016/MAX999 | LT1016CN8 | DIP-8 |

## 6.2 Signal Conditioning

| Qty | Description | Value |
| --- | --- | --- |
| 2 | Resistor | 10kΩ 1% |
| 1 | Resistor | 100Ω |
| 1 | Resistor | 2.2kΩ |
| 1 | Resistor | 3.3kΩ |
| 4 | Resistor | 330Ω |
| 1 | Capacitor | 100nF ceramic |
| 1 | Diode | BAT54 |
| 4 | LEDs | 3mm assorted |

# 7. Assembly Instructions

## 7.1 Preparation

- Gather all components from the BOM

- Test FE-5680A separately before integration

- Program Pico with CHRONOS-Rb firmware

- Prepare workspace with ESD protection

## 7.2 Assembly Steps

- Build 10MHz comparator circuit on perfboard

- Add bypass caps near comparator power pins

- Build 1PPS voltage divider

- Test each section before connecting to Pico

- Connect D-sub to FE-5680A

- Apply power and wait for lock (~5 min)

- Verify signals with oscilloscope

- Connect signal conditioning to Pico GPIOs

# 8. Firmware Overview

Build with Pico SDK 2.0+, CMake 3.13+, ARM GCC. Edit chronos_rb.h for WiFi credentials before building. Flash the UF2 file while holding BOOTSEL.

## 8.1 Module Overview

| Module | File | Function |
| --- | --- | --- |
| Main | main.c | Entry, init, loop |
| PPS | pps_capture.c | 1PPS timing |
| Freq | freq_counter.c | 10MHz measurement |
| Sync | rubidium_sync.c | State machine |
| Discipline | time_discipline.c | PI controller |
| NTP | ntp_server.c | NTPv4 server |
| PTP | ptp_server.c | IEEE 1588 |
| WiFi | wifi_manager.c | Connection mgmt |
| Web | web_interface.c | HTTP interface |

# 9. Using the Time Server

## 9.1 NTP Client Setup

Linux: Add NTP=<ip> to /etc/systemd/timesyncd.conf

Windows: w32tm /config /manualpeerlist:<ip> /syncfromflags:manual /update

macOS: System Preferences > Date & Time, or: sudo sntp -sS <ip>

## 9.2 Web Interface

Access http://<device-ip>/ for real-time status. JSON API at /api/status for monitoring integration.

# 10. Troubleshooting

| Symptom | Cause | Solution |
| --- | --- | --- |
| No 10MHz | Power issue | Check +15V/+5V |
| Won't lock | Warmup | Wait 5+ minutes |
| No 1PPS | Not locked | Wait for lock LED |
| WiFi fail | Credentials | Check config |
| High offset | Signal issue | Check circuits |
| NTP timeout | Firewall | Open UDP 123 |

# Appendix: Project Name

CHRONOS-Rb =

- Compact

- High-precision

- Rubidium

- Oscillator

- Network

- Operating

- System

The "-Rb" suffix denotes rubidium (chemical symbol Rb). Chronos (Χρόνος) is the Greek personification of time.

— End of Document —